| Model: |  |
| --- | --- |
| Serial Number: |  |
| Date: |  |
| Location: |  |

| # | Step | Result or Information |
| --- | --- | --- |
| IQ-1 | Unpack/Inspect the Incubator Shipping container and instrument intact and undamaged? Record information as applicable: Order Number: Ship Date: | YES NO N/A If NO, contact your carrier and Agilent or your supplier |
| IQ-2 | Check Required Components Components intact and undamaged? | YES NO N/A If NO, contact your carrier and Agilent or your supplier |
| IQ-3 | Check Optional Components Components intact and undamaged? If YES, list or attach documentation: | YES NO N/A If NO, contact your carrier and Agilent or your supplier |
| IQ-4 | Select an Appropriate Location Incubator located on level surface with temperature and humidity in acceptable range? | YES NO N/A |
| IQ-5 | Remove Shipping Hardware Shipping hardware removed and stored? Including: - shipping platform - robot shipping bracket - drawer shipping brackets Left-side reader installation steps performed? - feet and alignment plate installed | YES NO N/A YES NO N/A |
| IQ-6 | Set up the BioSpa Install internal components? Including: - Baffles - HEPA filter - Plate holders - Water pan (gas models only) - Close lid and install cover Install external components? Including: - Deck spacers - Gripper cover - Robot cover - Platform/cover - Guard rail | YES NO N/A YES NO N/A |
| IQ-7 | Install Companion Instruments Fluidics instrument installed on left-side platform? Detection instrument installed? | YES NO N/A YES NO N/A |
| IQ-8 | Install PC Software BioSpa software installed? Gen5 software/upgrade installed? LHC software/upgrade installed? | YES NO N/A YES NO N/A YES NO N/A |
| IQ-9 | Install Gas Tanks & Registers CO2 and N2 gas canisters installed? | YES NO N/A |
| IQ-10 | Connect Components to the Incubator Required components connected? Including: - CO2 & N2 tubing? - Platform/cover cable? - Liquid handler USB cable? - Power cable? | YES NO N/A |
| IQ-11 | Connect Host Computer to Incubator and Reader Companion reader and incubator connected to the host computer? If required: Record computer specifications (e.g., processor, operating system), or attach documentation. | YES NO N/A |
| IQ-12 | Connect Power Up to three power supplies/cords connected to power receptacles that provide voltage and current within the specified rating for the system? | YES NO N/A |
| IQ-13 | Configure Companion Instruments Reader configured in Gen5? Liquid handler configured in LHC? | YES NO N/A YES NO N/A |
|  | Establish Communication Communication with reader tested in Gen5? Communication with washer or dispenser tested in LHC? | YES NO N/A YES NO N/A |
|  | Create & Test Run Protocols Gen5 protocols tested? LHC protocols tested? | YES NO N/A YES NO N/A |
| IQ-14 | Configure Instruments for BioSpa BioSpa software configured to use liquid handler and/or reader? All self-tests successful? | YES NO N/A PASS FAIL N/A |
|  | Record BioSpa Software Info Model: Serial Number: Installation Version: PC Software Version: Motor Controller Basecode: Gas Controller Basecode: Temperature Controller Basecode: Gen5 Version: LHC Version: | N/A |
| IQ-15 | Align Instruments Gripper aligned to deliver and retrieve microplates to companion instruments? | YES NO N/A |
| IQ-16 | Create and Run a Test Session Define a Reader-Imager session? Define a Washer-Dispenser session? Gas-flow test pass? Session completed successfully? Including: - Temperature at set point - CO2/N2 at set point - Humidity at expected level | YES NO N/A YES NO N/A YES NO N/A YES NO N/A |
|  | Review Safety Information Information pertaining to Warnings, Hazards, Precautions, and Safety Symbols reviewed (or will be reviewed, as a part of training)? | YES NO N/A |
|  | Product registered with Agilent? | YES NO N/A |

Comments:

| Installation Qualification conducted by: | Date: |
| --- | --- |

PRINT Name:

| If required, reviewed/approved by: N/A | Date: |
| --- | --- |

PRINT Name: